## Slide 1
Crisp-BP: Continuous Wrist PPG-based Blood Pressure Measurement
Presenter: Yetong Cao
Yetong Cao∗ Huijie Chen† Fan Li∗ Yu Wang‡
*School of Computer Science, Beijing Institute of Technology, China
† School of Computer Science, Beijing University of Technology, China
‡ Department of Computer and Information Sciences, Temple University, USA
MobiCom 2021

## Slide 2: Background
Hypertension: blood pressure is too high
Global health issue: affect 1 in 5 people
Silent killer: no symptoms
Damage the body
Blood pressure is the key signal in the prevention and diagnosis of hypertension.

## Slide 3: Current BP Monitoring Solutions
Related Solutions
Arm-cuff BP Monitoring
Work discontinuously
https://www.youtube.com/watch?v=N2GZCfJ5lK8

## Slide 4: Current BP Monitoring Solutions
Related Solutions
Pulse Wave Velocity (PWV)-based BP Monitoring
Require user effort

## Slide 5: Our Solution
PPG sensor
160
90
Crisp-BP
Wrist-worn devices
Low cost
Comfortable
Convenient
Support long-term monitoring

## Slide 6: Our Solution
Reflected Wave Transit Time (RWTT)=

## Slide 7: Challenges
Capillaries
Arteries
Veins
160
90
Crisp-BP
PPG LED
Photodetector
Infrared PPG
Capillary Pulse
Arterial Pulse
How to obtain clean arterial pulse?

## Slide 8: Challenges
Capillary
Artery
PPG Sensor
Low Pressure
Proper Pressure
High Pressure
How to address contact pressure artifact?
Time
Pressure Level
AC Component

## Slide 9: Challenges
How to enable accurate and user-independent ABP monitoring?

## Slide 10: Measuring RWTT using Acceleration PPG

## Slide 11: System Overview
Yes
No
Adjust the writ band
Address challenge #1:
Help users maintain appropriate contact pressure
Address challenge #2:
Extract clean arterial pulse
Address challenge #3:
Realize accurate, user-independent ABP estimation

## Slide 12: Device State Identification
Contact Pressure Estimation
  Extract two unique features
  Estimate contact pressure using Least-Squares Support Vector Machine (LSSVM)

## Slide 13: Device State Identification
Contact Pressure Qualification
  Determine the Optimal Contact Pressure (OCP)
  Guide the user to adjust the wristband to achieve OCP
OCP

## Slide 14: Arterial Pulse Profiling
Noise Reduction
  Bandpass filtering: 4th-order Butterworth, [0.5-8] Hz
  Remove outlier pulses: percentage change method (Malik,1990)
Arterial Pulse Extraction
Capillaries
Arteries
Veins
Infrared LED
Photodetector
Infrared PPG
(Infrared PPG)
Capillary Pulse
(Green PPG)
Arterial Pulse
Green LED
remove

## Slide 15: Arterial Pulse Profiling
Capillaries
Arteries
Veins
Infrared LED
Photodetector
Green LED
Optical Density Modeling：

## Slide 16: Continuous ABP Monitoring
Requiring sufficient training data

## Slide 17: Continuous ABP Monitoring
Boosting the ABP Monitoring
weights & bias

## Slide 18: Evaluation
Setup
  PPG-based prototype
Data Collection
  35 participants (17 male, 18 female, aging between 19 to 50)
  10 sessions
  Continuous monitoring experiment
  Various PPG sensors, sampling frequencies, contact pressures, and network structures
Ground Truth
  Omron U30 device
  Standard validation procedure
  prototype
  Arm-cuff BP monitoring device

## Slide 19: Evaluation
Metrics
  Mean error
  Standard deviation
  Sample Pearson’s correlation coefficient

## Slide 20: Overall Performance
Mean error：
  DBP 0.86 mmHg
  SBP 1.67 mmHg
Bland–Altman diagram of DBP and SBP
Standard deviation
  DBP 6.55 mmHg
  SBP 7.31 mmHg

## Slide 21: Overall Performance
Sample Pearson’s correlation coefficient
  DBP 0.88
  SBP 0.82
Correlation diagram of DBP and SBP

## Slide 22: Overall Performance
 |  | ME (mmHg) | STD (mmHg)
Crisp-BP | DBP | 0.86 | 6.55
 | SBP | 1.67 | 7.31
AAMI | DBP & SBP | 5 | 8
 |  | ME (mmHg) | STD (mmHg)
Crisp-BP | DBP | 0.86 | 6.55
 | SBP | 1.67 | 7.31
AAMI | DBP & SBP |  | 
 |  | Cumulative Error Percentage |  | 
 |  | 5 mmHg | 10 mmHg | 15 mmHg
Crisp-BP | DBP | 76.73% | 91.43% | 97.96%
 | SBP | 65.66% | 87.17% | 96.23%
BHS | Grade A | 60% | 85% | 95%
 | Grade B | 50% | 75% | 90%
 | Grade C | 40% | 65% | 85%
 |  | Cumulative Error Percentage |  | 
Crisp-BP | DBP | 76.73% | 91.43% | 97.96%
 | SBP | 65.66% | 87.17% | 96.23%
BHS | Grade A | 60% | 85% | 95%
 | Grade B | 50% | 75% | 90%
 | Grade C | 40% | 65% | 85%
Association for the Advancement of Medical Instruments
Britain Hypertension Society

## Slide 23: Comparison with Typical Related Works
Items | Glabella | SeismoWatch | Seismo | Naptics | eBP | Crisp-BP
Method | PTT | PTT | PTT | PTT | CPA | RWTT
Subject | 4 | 13 | 9 | 6 | 35 | 35
Device | Glasses | Wristband | Phone | Shorts | In-ear | Wristband
Auxiliary Tools | - | Accel | Accel | Accel,elec,oxi | Balloon | -
Passive Sensing |  |  |  |  |  | 
Scenario | Daytime | Instant | Instant | Night | Instant | Daytime & night
Comfort Level | High | Middle | Middle | Middle | Middle | High
SBP ME/STD (mmHg) | 10 | 4.8 | - | - | 1.8/7.2 | 1.67/7.31
DBP ME/STD (mmHg) | - | 2.9 | 1.3/6.9 | - | -3.1/7.9 | 0.86/6.55
Items | Glabella | SeismoWatch | Seismo | Naptics | eBP | Crisp-BP
Method | PTT | PTT | PTT | PTT | CPA | RWTT
Subject | 4 | 13 | 9 | 6 | 35 | 35
Device | Glasses | Wristband | Phone | Shorts | In-ear | Wristband
Auxiliary Tools | - | Accel | Accel | Accel,elec,oxi | Balloon | -
Passive Sensing |  |  |  |  |  | 
Scenario | Daytime | Instant | Instant | Night | Instant | Daytime & night
Comfort Level | High | Middle | Middle | Middle | Middle | High
SBP ME/STD (mmHg) |  | 4.8 | - | - | 1.8/7.2 | 1.67/7.31
DBP ME/STD (mmHg) | - | 2.9 | 1.3/6.9 | - | -3.1/7.9 | 0.86/6.55
PTT, CPA stand for pulse transit times, change of pulse amplitude, respectively. Accel, elec, and oxi represent accelerometer, electrodes, and oximeter, respectively

## Slide 24: 24-Hour Performance
24-hour performance
Mean error：
  DBP <3.2 mmHg
  SBP <3.2 mmHg

## Slide 25: Conclusion
We propose an effective system to monitor ABP.
We design a set of techniques to ensure accurate ABP measurements.
Evaluation Results are promising.
More details？
Yetong Cao, Huijie Chen, Fan Li, Yu Wang,
Crisp-BP: Continuous Wrist PPG-based Blood Pressure Measurement,
In MobiCom 2021, New Orleans, United States, January 31 - February 04 2022

## Slide 26
Thanks!
Thanks!
Thanks!
&
&
&
Questions
Questions
Questions
Yetong Cao
yetongcao@bit.edu.cn