## Slide 1
hBP-Fi: Contactless Blood Pressure Monitoring via Deep-Analyzed Hemodynamics
Yetong Cao∗, Shujie Zhang†, Fan Li∗, Zhe Chen‡, Jun Luo†
*Beijing Institute of Technology, China
†Nanyang Technological University, Singapore
‡Fudan University, China
INFOCOM 2024

## Slide 2: Background
Hypertension: blood pressure is too high
Global health issue: affect 1 in 5 people
Silent killer: no symptoms
Damage the body
Blood pressure is the key signal in the prevention and diagnosis of hypertension.

## Slide 3: Current BP Monitoring Solutions
Contact-based BP Monitoring
Arm-cuff solutions
Pulse wave velocity (PWV)-based solutions
Cuff-based wearable solutions
Uncomfortable
Unsuitable for individuals with sensitive or injured skin
Strict requirements on wearing position and pressure

## Slide 4: Current BP Monitoring Solutions
Contactless BP Monitoring: PWV-based solution is impractical!
Contactless Blood Pressure Monitoring with mmWave Radar, IEEE GLOBECOM 22
airBP: Monitor Your Blood Pressure with Millimeter-Wave in the Air, ACM TIOT 23
mmBP: Contact-Free Millimetre-Wave Radar Based Approach to Blood Pressure Measurement, SenSys 22
Trust issue due to lack of a physiological basis

## Slide 5: Our Solution: hBP-Fi
Beam Scan
Compliance Changes
Explainable DL
BP
BP-Encoded Hemodynamics
hBP-Fi associates advanced RF sensing and DL inference with a sound physiological basis.
Varying Pulse Waveforms
Related to BP

## Slide 6: Our Solution: hBP-Fi
Beam Scan
Compliance Changes
Varying Pulse Waveforms
Explainable DL
BP
BP-Encoded Hemodynamics
How RF signals respond to hemodynamics?
How to capture varying pulse waveforms along the arm?
How to estimate BP explainablely？

## Slide 7: System Overview
We introduce BTF for correlating RF-captured pulse waveforms with BP.
We develop an explainable DL pipeline for BP inference.
We propose a continuous beam scan RF sensing scheme.

## Slide 8: Hemodynamics Modeling via BTF
1
Blood pressure value

## Slide 9: Hemodynamics Modeling via BTF
1
Blood pressure value

## Slide 10: RF Pulse Waveform Collection
2
How can we achieve a clearer separation of distinct areas?
Magnify the difference of the RF reflections from two neighboring sites!

## Slide 11: RF Pulse Waveform Collection
Super-Resolution Beam Scan
2
Null-steering
Pulse at site #1
Pulse at site #2

## Slide 12: RF Pulse Waveform Collection
Super-Resolution Beam Scan
Pulse Waveform Extraction
2
FMCW
mmWave Radar
Differential
RF reflections
IF signal
Radar data tensor
Phase variations
FFT

## Slide 13: RF Pulse Waveform Collection
Super-Resolution Beam Scan
Pulse Waveform Extraction
2
BP results are not accurate enough!
BTF model fails to address influence of temperature and time of day on BP in real-world environments .

## Slide 14: BP Inference
3
BTF model
DL Technique
Structure & Searching Space

## Slide 15: BP Inference
Explainable DL Pipeline
3

## Slide 16: Evaluation
Setup
  TI AWR1843BOOST mmWave radar, 77GHz, 4G bandwidth, 3Tx, 4Rx
  DCA 1000EVM data capture board, 200 frames/second
Data Collection
  35 participants (22 male, 13 female, aging between 20 to 54)
  Minimal body movement, hold breath
  Varying environments
  Varying radar placements
Ground Truth
  Arm-cuff BP monitor
  Standard procedure

## Slide 17: Evaluation
Metrics
  Mean error
  Standard deviation of mean error

## Slide 18: Overall Performance
Mean error：
  SBP -2.05 mmHg
  DBP 1.99 mmHg
Standard deviation
  SBP 6.83 mmHg
  DBP 6.30 mmHg

## Slide 19: Overall Performance
 |  | ME (mmHg) | STD (mmHg)
hBP-Fi | DBP | 1.99 | 6.30
 | SBP | -2.05 | 6.83
AAMI | DBP & SBP | 5 | 8
 |  | ME (mmHg) | STD (mmHg)
hBP-Fi | DBP | 1.99 | 6.30
 | SBP | -2.05 | 6.83
AAMI | DBP & SBP |  | 
 |  | Cumulative Error Percentage |  | 
 |  | 5 mmHg | 10 mmHg | 15 mmHg
hBP-Fi | DBP | 62.57% | 87.82% | 98.93%
 | SBP | 61.50% | 86.63% | 97.65%
BHS | Grade A | 60% | 85% | 95%
 | Grade B | 50% | 75% | 90%
 | Grade C | 40% | 65% | 85%
 |  | Cumulative Error Percentage |  | 
hBP-Fi | DBP | 62.57% | 87.82% | 98.93%
 | SBP | 61.50% | 86.63% | 97.65%
BHS | Grade A | 60% | 85% | 95%
 | Grade B | 50% | 75% | 90%
 | Grade C | 40% | 65% | 85%
Association for the Advancement of Medical Instruments
Britain Hypertension Society

## Slide 20: Robustness
 |  | 30cm | 50cm | 70cm | 90cm
SBP | 0 | -2.05±6.76 | -2.04±6.72 | -2.35±7.21 | -3.47±7.12
 | 15 | -2.00±6.74 | -2.03±6.77 | -2.15±7.36 | -3.46± 7.61
 | 30 | -2.02±6.81 | -2.07±6.91 | -2.48±7.11 | -3.46±7.79
 | 45 | -2.06±6.92 | -2.05±6.85 | -2.10±7.20 | -3.48±7.77
DBP | 0 | 1.99±6.31 | 2.03±6.82 | 2.68±7.47 | 4.33±7.22
 | 15 | 2.07±6.30 | 2.04±6.51 | 2.69±6.63 | 3.34±7.30
 | 30 | 2.09±6.35 | 2.09±6.53 | 2.68±6.78 | 3.29±7.72
 | 45 | 2.03±6.34 | 2.06±7.72 | 2.69±7.28 | 3.36±7.87
hBP-Fi is practical in real-world scenarios!
Varying environments
Varying radar placements

## Slide 21: Conclusion
Target
Systems
Challenges
Solutions
Results
Contactless BP Monitoring
hBP-Fi
A sound physiological basis
Accurate BP results
BTF Model
Beam Scan
Explainable DL
Unclear
Relationship
Low Quality
Inaccurate

## Slide 22
Thanks!
Thanks!
Thanks!
Q&A
Q&A
Q&A
Yetong Cao
yetong.cao@ntu.edu.sg
Google Scholar Page
Jun Luo, IEEE Fellow
junluo@ntu.edu.sg
Google Scholar Page